Lumencor web site http://lumencor.com/products/spectra-x-light-engine/

Available filters:

- Violet: 395/25

- Blue: 440/20

- Cyan: 470/24

- Red: 640/30

- Teal: 510/25

- Green/Yellow: 550/15 or 575/27 (exchange filters for optimal excitation)

Troubleshoot:

- When light remains on all the time, check cable connection in the lumencor configuration settings (see figure)

- When light is not tuneable by software miscommunication via the COM port likely occurs (in: >Device select SpectraX > settings => set COM port to 2)

mCherry imaging:

575/25  filter should be in the SpectraX light house (extra filter is placed below the cover in the light house)

Place in the Green/Yellow  (GY) slit grip facing G/Y marking (does not fit otherwise)

- Select in config : 575 nm, check connections.

- Use CFP/YFP/mCherry filter block in lower turret

- Image as usual (use ND4 and 8 filters to tune the LED light at lower levels)