# Workflow Acquisition in LAS-X: Quickstart

Contents

Prepare	1

File Save Preparation and Focus on cells	1

Setting the detection filters (Image Compass)	1

Tweak Acquisition Scanning parameters:	3

Using Navigator: Positions - Carrier – ROI Sections	3

Lightning Mode: Automate parameter setting to perform live deconvolution	5

Tau Separation: lifetime () isolation or gating ( >1 ns)	6

……………………………………………………………………………………………………………

## Prepare

- Make notes of settings, rank your nicest images (link for a CCI-prepared template)

- Configuration tab: Bit depth 8<>16 bit, Glow Over, HyD’s in Count mode recommended

- Is optimal pixel size required for the experiment? Choose suited Objective and adjust pixel size to meet Nyquist criteria only when you need it

- Decide if, and which detector setting uses the transmitted light detection

- Know the dyes in your samples (spectral, lifetimes if available, stability)

  - Excitation/Emission spectra can be shown/uploaded in LAS-X dye database (Configuration tab)

## File Save Preparation and Focus on cells

- Save Preparation: <PROJECTS> presets File Naming and file location

  - <.xlef> = lif+ extended metadata

  - In <New Project Defaults> set Project name: <YYYYMMDD_#UID-RefID (link to LabJournal)>

  - set <User defined name> #UID-RefID_TXT-describes-experiment_Channel-desciption

  - NO special characters (!@#$%^&*(){}?|\/=][><~`”;:) or SPACES, < 256 characters (with folder names)

  - <Autosave> to save multidimension data during acquisition (tick <OFF> while tweaking)

  - File Location: subfolders on E:\ImageDATA\<USER>\<DATE: YYYMMDD>

  - Make a <New Project> when switching samples (LAS-X add counters, Well-/ROI-name), re-save list frequently.

- Focus on cells

  - Transmission images usually work fine, low Mag 5x/10x, find a cs/well edge and move to cells, use a simple method or view by eye

## Setting the detection filters (Image Compass)

- Start in Acquisition Mode (in Viewer: Split Channel / Activate Intensity Histogram)

- Optimize detector settings: Reuse/Create a detection sequence from file metadata or saved setting file => CHECK before use (LabJournal notes - link)

  - Hardware: Objective, filter ranges, laser (, power), HyD (type, mode, Gain, AI denoise),

  - Acq Settings: Pixel Size (Format & Zoom), Pix Dwell time (Speed-Galvo/Res, Bi-Dir), Pinhole, Acc/Avg, Z step, Z  range, t interval-duration

- Search Dye, select and drag it onto the preferred HyD Detector (or drag onto empty space for automated setting), repeat for every dye

Figure1: HyD properties

Choose optimal HyD Type for Dye (order S1>X2>S3>X4>R5, start at shortest wavelength)

- Choose sequential mode to record dye (sets)

  - <+> Add a detector group, Ms(R) to reset a group

  - Line recommended, Frame allows different Acc or Avg per group

  - <Uncheck> to temporarily skip its acquisition

  - Avoid Dye Assistant, it resets filter ranges and does not include the HyD-R detector of the system (use it to get dye combination hints, do not <Apply>)

- Choose mode of the HyDs (analog/digital/photon count (FLIM)/reflection)

- Set Emission Filter Slider range

  - Fit emission spectra (shallow =>  Crosstalk &  Signal), tweak Yield/Crosstalk by shifting the range

  - Ms(Clk)2 on slider to type -limits or Clk-n-Drag ends

  - No overlap with laser line (>8nm away) & AFC aborts when a filter 750 nm

- Set Excitation Laser

  - Fit excitation spectra – reduce Crosstalk by shifting from optimal, use PanelBox or Ms(Clk)2 the line to type

  - Inspect bleed/Crosstalk by varying Laser Intensity for the ‘donor’ channel, Intensity of the ‘acceptor’ Channel should not change, best evaluated in Viewer with Split channel & Histogram activated

  - Interactive elements open different setting GUIs (see numbered items in figure 2)

- Choose group that generates the Transmission channel (TransPMT >> <ON>, set its Gain/Offset)

- Set LUTs or channel coloring and save the setting in your data folder

  - Recommend to choose color blind-friendly sets that work well for presentation via beamers etc.)

  - Blue> Cyan, Red> Orange/Yellow, NIR> Magenta, TM > Grey; Explore Inverted for weak/sparse labels

Figure 2: Interactive GUI elements in the Image Compass

Use CCI form (on desktop or CCI website) dedicated to structure documentation of settings, save as: Setting-UID in your data folder before filling the form. Setting comments, for example Z range, Z step, T Interval-duration, image-Q rank

## Tweak Acquisition Scanning parameters:

#### Start/Stop scanning

  - <FAST Live> to scan quickly (fast setting tweakable, it scans only the activate -white marked- Detector Group)

  - <Live> scans all channels as set (starts by Panelbox <Right Button>)

  - <Capture> takes image channels and omits other dimensions

  - <START> records all set dimensions

  - Compare images in <Viewer> by <+> to add another Project Image, linking synchs Zoom/Intensity changes, <X> to remove a parallel view

#### Acquisition Parameter setting

- Laser Intensity (use the Panel Box <Smart Intensity> or Ms(Clk)2 line to type)

- Scan Acquisition mode XY

  - Do not oversaturate: Bit type intensity maxima (8b>256, 16b> 65536), use Glow-Over LUT in viewer to evaluate

  - <Zoom in> prompts to make a ROI and Zoom into it

Figure3: Acquisition Scanner Setting

<Bidirectional X> scan @ double speed but striping/doubling artefacts can occur (changing with zoom, speed>700, format change etc.), it needs manual Phase-X correction

  - <Format> & <Zoom Factor>  define XY Pixel Size (diffraction limit or not)

  - <Speed> defines Pixel Dwell Time: slower => signal , the <Resonance Mode> locked high speed: needs higher laser powers, and avg/acc to improve S/N

  - <Pinhole> defines along with  Optical section & Z step size (choose Airy 1 for the most important channel em)

  - <Rotation> orientate sample to optimally fit field

- Detector setting – signal pooling

  - <Gain> digitally enhances HyD signal – beware to go above 75 (not available in counting mode). TransPMT  is more forgiving in its setting

  - <Offset> only available for TransPMT (use to tweak lowest signals), stay close to 0

  - <Avg> scans line/frame and retains the avg intensity: good for Gain, i.e. non-count modes

  - <Acc> accumulates line/frame intensities: perfect for count mode HyD

- Define Multidimensional imaging (Z, t, position, etc.=> Select Detection mode)

- Use <Navigator> to <Align> and use preset Carriers or define ROIs/positions.

- Check focus control (AFC, Autofocus in noise, make Focus Map in Navigator)

- Check timing

## Using Navigator: Positions - Carrier – ROI Sections

Many of the acquisition GUI is available in the <Navigator> as well in grouped GUIs. Beam Path= image Compass, Scan Mode = XY scan parameters, Z stack, Time, positions GUIs are activated in the <Toolbar>. Generally, preparing imaging is easiest done in the acquisition tab while overview/position/tile/ROI functions are done in the navigator.

  - Ms(Clk)2 move to position and makes <Preview> image, Ms(Wheel) to Zoom.

  - <AFC> performs quite well in <Spiral Scan>, Ms(R) to select variants of the spiral scan

  - Avoid LUT pileup of all (tiles and) previews in the images listing by unchecking or deleting unnecessary previews/Tiles. Setting LUT for the active FOV is quite difficult otherwise.

  - <Split Channel View>  for active position FOV, a GUI offers options to overlay. Switch back to <Stage Overview> to resume tile actions

  - <Task List> Check/Unchecking ROIs/wells can be per group as well. <START> will only record checked <Tasks>

  - <Save Overview…> saves the checked pre/over-views in <Projects>

  - <Mosaic Merge> generates a single large file in <Projects>

Figure 4: Navigator

#### Focus MAP

- Make Spiral overview (low mag)

- Select ROI with drawing tools

- <  > keep Drawing Tool active

- <+F> to set focus XY position (remove by Select> Delete)

- <+AF> to use in Timelapse / Stage (set <..on Demand> – everty xth in time / position)

- <Focus Map…> to activate dialog and list XY

- <Set Z> Manual /  Autofocus for active or all point(s) => Values Z/AFC are shown

  - OPTION: Shift Entire MAP in Z by defined offset (workaround to do asymmetric stacking)

  - CONTEXT Menu: Subtract Tool / Group Scan Regions / Paste Regions in Chambers etc.

## Lightning Mode: Automate parameter setting to perform live deconvolution

Figure 5: Lightning mode GUI

Autoparameter setting (unchain to adjust setting)

- LAS-X gets immersion media refractive index, estimates of imaging depth aberrations and psf’s for the channels from metadata

- Adaptive  = generally better than Global

- Set the mounting media correctly (see table below)

Deconvolution is also available in the Process tab

Table 1: Lightning parameters

| Auto parameter | Speed | <> | Resolution |
| --- | --- | --- | --- |
| Pinhole (AU) | ~ 1 | <> | Up to 0.3 AU (optimal 2 better than Abbe) |
| Pixel/ Z-step | .. | <> | Nyquist for given Pinhole |
| Speed | .. | <> | Decreases to improve light collection |
| Line Avg | ~1 | <> | Increases for better S/N |
| HyD Gain | Up to you | <> | Up to you |
| Laser % | Up to you | <> | Up to you |

Table 2: Immersion media Refractive Indices

| Name | Refractive index |
| --- | --- |
| VectaShield Vibrance | 1.47 (cured) |
| VectaShield Plus | 1.45 |
| VectaShield Harset | 1.46 (cured) |
| VectaShield | 1.45 |
| SlowFade Gold | 1.42 |
| SlowFade Glass | 1.52 |
| SlowFade Diamond | 1.42 |
| ProLong Gold | 1.46 |
| ProLong Glass | 1.51 |
| Prolong Diamond | 1.46 |
| Mowiol | 1.42 |
| Fluoromount G | 1.42 |
| Entellan | 1.49 |
| DPX | 1.52 |
| Canada Balsam | 1.52 |

Table 3: Laser Frequency  to interval conversion

| Frequency (MHz) | Interval (ns) |
| --- | --- |
| 80 | 12.5 |
| 40 | 25 |
| 20 | 50 |
| 10 | 100 |
| 5 | 200 |
| 2.5 | 400 |

Tau Separation: lifetime () isolation or gating ( >1 ns)

Average Photon arrival time per pixel => assign Pix to an avg  class

#### Prerequisites:

  - WLL (OPTION: notch filtered reference lines) / TauContrast <ON> / HyD in count mode

  - WLL pulse frequency should accommodate 3-4x the expected longest lifetime

  - Maximal Photon count rate should be Pulse rate  0.1 / laser pulse (Pixel Int Histo)/ 0.37 with Photon count filter

  - Tau output images are kept as separate channels in <Projects>

- TauContrast (analyse lifetime components in an emission block) > Intensity

  - OUTPUT: split viewer channel of emission block: Intensity channel and Lifetime channel (Spectral LUT-encoded , => spot peaking colors to identify abundant  subsets )

- TauScan (set nr of images to extract  -spieces)

  - OUTPUT: nr series with -range/nr, => spot the peak intensities in the image range

- GateScan (nr of preset Gates to estimate -species): <+> to add onther channel and slider set to define bounds

  - OUTPUT: series of defined -”bins” => spot differences in the extracted images

- TauGating:  Set  range based on Tau/Gate-Scan to put in a channel (unchecked range(s) will not be recorded => omit reflections, autofluorescence)

  - OUTPUT: Tau Images with selected  characteristics (and omitted  species from the recorded images)

- TauSeparation (use  estimates from Tau/Gate-Scan, <+> and define in slider, or allow algorithm to extract them, <Find>)

  - OUTPUT: Tau Images with defined  characteristics